## Gauge
|  | Test Date | 2014-07-06 |  | Date of Factory Calibration | 2014-07-06
|  | Gauge Model | DM32 |  | Technician | John Smith
|  | DM32 Serial Number | 402102 |  | Company | Company Name
|  | Retrotec Fan Series | 6000
|  |  | Channel A pressure |  | Channel B pressure |  | Difference from Average
Step 1: | Connect Yellow | tube from: | Blue port | to | Yellow port
|  |  | 25 |  | 25.6 | ⇨ | 0.3 Pa | PASS | 1.186% | 0.3
Step 2: | Connect Yellow | tube from: | Red port | to | Green port
|  |  | 242 |  | -243 | ⇨ | 0.21 % | PASS | 0.2062% | 0.5
Step 3: | Connect Red | tube from: | Red port | to | Green port
|  |  | 64.4 |  | -67.7 | ⇨ | 2.5 % | FAIL | 2.498% | 1.65
Step 4: |  | No tubes connected: | Channel A |  | Channel B
|  |  | 0.1 |  | 0.2 | ⇨ | 0.1 Pa | PASS | 33.33% | 0.05
[8 empty] | This gauge FAILS the field calibration check | FAIL
If the gauge FAILS, repeat Steps 1-4
If the gauge still FAILS, contact support@retrotec.com
Field Calibration Check for Gauge
This procedure checks the calibration of each channel against the other channel in both positive and negative directions. It also checks all tubes used for leaks and blockages. Pressure will drop slightly each time the gauge auto-zeros but if it drops rapidly try another tube. If it drops to zero in 20 seconds there may be an internal leak in the gauge. If the pressure is more than 10% different between channels, try another tube. In both cases, the tubing should be have no water drops in it; nor should the gauge connections.
If tubing is causing the problem, replace the tube or dry it out by swinging it around. If the gauge is still more than 2% different between channels, it should be returned for calibration or checked against another gauge if one is available. Make sure you know which component is faulty before returning it for calibration.
Instructions:
□ Place the gauge and tubing on a stable surface where they will not move. Away from heat sources such as hands, heaters or the Sun and away from cold sources. They must be indoors at the same temperature for 30 minutes.
□ Turn on the gauge.
□ Set [Time Average] to 5 seconds in [Settings].
□ Tap [Channel B] and select "Pressure: Pa".
□ Perform Step 1 below, and wait 60 seconds.
□ Tap Channel A to activate [HOLD]. Enter Channel A and Channel B pressures, then tap Channel A to release [HOLD].
□ Perform Step 2 below, and wait 15 seconds.
□ Tap Channel A to activate [HOLD]. Enter Channel A and Channel B pressures, then tap Channel A to release [HOLD].
□ Perform Step 3 below, and wait 15 seconds.
□ Tap Channel A to activate [HOLD]. Enter Channel A and Channel B pressures, then tap Channel A to release [HOLD].
If Channel A and Channel B readings are within 1% and don’t drop rapidly, the gauge is in calibration and the tube is not blocked or leaking.
PASS/FAIL: ±1% of reading or ±0.5 Pa
Check your gauge and tubes before every important test to ensure they will give correct results.
Fill out the test form to verify the Field Calibration Check has been performed at time intervals required by the Standard.
Note on D30: Date format: YYYY-MM-DD YYYY=Year MM=Month DD=Day